PERFORMANCE OF THE AD590 TEMPERATURE TRANSDUCER AT LOW TEMPERATURE 

Test Report

Richard Patterson

NASA Glenn Research Center

&
Scott Gerber

ZIN Technologies 
&

Ahmad Hammoud

QSS Group, Inc.
NASA Glenn Research Center

Cleveland, Ohio

August 10, 2001
Performance Of The AD590 Temperature Transducer At Low Temperature
Device Description


The AD590 is a two-terminal integrated circuit (voltage in/current out) temperature transducer that produces an output current proportional to absolute temperature.  In addition to temperature measurement, this device can be used in applications that include flow rate measurements, correction of discrete components, temperature compensation, biasing proportional to absolute temperature, level detection of fluids and anemometry.

Effort Objective


Temperature transducers are commonly used in electronic systems in almost all of NASA missions.  The simple design of the AD590, which requires no support circuitry, makes it particularly useful in remote sensing applications [1].  Unlike other conventional transducers, the AD590 does not require any linearization circuitry, precision voltage amplifiers, resistance measurement circuit, or cold junction compensation.  Its simple design and high accuracy render it a good candidate for use in many of NASA space exploration missions.

Device Specifications


The AD590 temperature transducer is made by Analog Devices, Inc. and Harris Corporation. The device comes in different grades with ceramic package or metal can options. The specifications of some of these packages are listed in Table I.

Table I.  Some of the manufacturer specifications of the devices tested [2,3].

	
	AD590LF
	AD590JF
	AD590JH

	Input Voltage (V)
	+4 to +30
	+4 to +30
	+4 to +30

	Linear Current Output ((A/K)
	1
	1
	1

	Temperature Range ((C)
	-55 to +150
	-55 to +150
	-55 to +150

	Maximum Nonlinearity ((C)
	(0.4
	(1.5
	(1.5

	Maximum Absolute Error ((C)
	(3
	(10
	(10

	Maximum Calibration Error @R.T. ((C)
	(1
	(5
	(5

	Maximum Repeatability ((C)
	(0.1
	(0.1
	(0.1

	Package
	Ceramic
	Ceramic
	Metal


Test Procedure

Three different versions of the AD590 were evaluated in terms of output response and linearity as a function of temperature.  A total of 61 devices were characterized in the temperature range of 293K  (+20 (C) to 80K (-193 (C). The data were initially gathered at room temperature, then at successive test temperatures down to 80K, and again at room temperature after low temperature was completed. At a given test temperature, the device under test was allowed to soak for 20 minutes so that thermal equilibrium was reached before any measurement was made. To investigate the effect of low temperature exposure on the packaging, devices with ceramic packaging and metal cans were used in this work.  Each device was characterized in its output current, which is proportional to absolute temperature, as a function of temperature and its input or supply voltage. After the last measurement was taken at the lowest temperature, i.e. ‑190 (C, data was then obtained at room temperature.

 Test Results

A total of 61 devices were evaluated in this work. These comprised of ceramic-packaged (LF and JF versions) and metal-canned (JH version) units. While the first batch (six) of the LF‑series was about seven years old, the rest of the devices (55) were recently acquired. An overall summary on the performance of the devices tested is shown in Table II. 

Table II. Summary of performance of the various AD590 devices.

	Version
	Package
	Quantity
	Manufacturer
	Performance with Temperature

	LF
	Ceramic
	6
	Analog Devices

(old)
	5 devices showed good performance down to 90 - 80K

1 device had good performance down to only 144K

	
	
	6
	Analog Devices


	5 devices showed good performance down to 170 - 140K

1 device exhibited deviation in performance below 220K

	JF
	Ceramic
	30
	Analog Devices 
	9 devices exhibited good performance down to 120 - 80K

19 devices showed good performance down to (130K 

2 devices started to drift in performance below 170K

	JH
	Metal
	10
	Analog Devices
	All devices performed well down to only 120K

	
	
	9
	Harris
	All devices exhibited poor performance between 220-170K


It can be clearly seen from Table II that a wide variation in performance was exhibited by the various versions of the AD590 temperature transducer.  In fact, discrepancies in output behavior with temperature were even observed for devices from the same series. For example, the first batch (old six) of the ceramic LF-series exhibited, in general, much better performance with temperature than those acquired recently.  It is not readily clear why such unexpected outcome is observed.  However, it can be postulated that different production lots and possibly new manufacturing processes or design modifications might have contributed to this behavior.

Most of the JF-series devices exhibited good performance down to about 140 to 130K.  While nine of the remaining devices performed relatively well with temperature down to about 120 to 80K, two devices exhibited good operation only down to 170K.  It is important to note that all of the devices were able to generate some output beyond the aforementioned good low temperature cut-offs.  The produced output, however, did not reflect an accurate representation of the temperature to be measured.

The metal-can packaged devices were obtained from two manufacturers. The ten Analog Devices units displayed good and consistent behavior from room temperature down to 120K.  The Harris devices, nine of them, however, exhibited behavior that was different from one device to another from the same batch.  While some of these devices operated quite well down to 170K, others exhibited irregularity in their output at temperatures as low as 220K.

A detailed listing of the devices tested along with their corresponding good low temperature limit, below which deviation in the output was observed, is given in Table III.

Table III. Approximate working low temperature of devices tested.

	Device Type
	Device #.
	Temp. (K)
	Device Type
	Device #.
	Temp. (K)

	AD590LF(old)
	1
	80
	AD590JF
	4
	123

	
	2
	80
	
	5
	123

	
	3
	144
	
	6
	102

	
	29
	80
	
	7
	144

	
	30
	80
	
	8
	144

	
	31
	90
	
	9
	123

	AD590LF(new)
	56
	170
	
	32
	100

	
	57
	220
	
	33
	110

	
	58
	150
	
	34
	170

	
	59
	140
	
	35
	140

	
	60
	150
	
	36
	130

	
	61
	140
	
	37
	130

	AD590JH
	10
	120
	
	38
	130

	Analog
	11
	120
	
	39
	140

	
	12
	120
	
	40
	80

	
	13
	120
	
	41
	140

	
	14
	130
	
	42
	130

	
	15
	120
	
	43
	130

	
	16
	120
	
	44
	130

	
	17
	120
	
	45
	120

	
	18
	120
	
	46
	130

	
	19
	120
	
	47
	130

	AD590JH
	20
	220
	
	48
	130

	Harris
	21
	220
	
	49
	120

	
	22
	170
	
	50
	170

	
	23
	220
	
	51
	140

	
	24
	220
	
	52
	130

	
	25
	170
	
	53
	130

	
	26
	220
	
	54
	140

	
	27
	170
	
	55
	130

	
	28
	220
	
	
	



Graphical representations of the data obtained on each device are followed in the same sequence as were presented in Table II. The parameters depicted in each figure include the supply input voltage (a test variable) and the specific test temperatures at which measurements were made. The corresponding output current, which was measured at various temperatures, is plotted along the coordinate. This measurement can also be read as temperature in Kelvins since the AD590 temperature transducer provides an output of 1(A/K.

[image: image1.wmf]0

5

10

15

20

25

30

Supply Voltage (V)

AD590LF Dev #1

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

5

10

15

20

25

30

Supply Voltage (V)

AD590LF Dev #2

0

50

100

150

200

250

300

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

5

10

15

20

25

30

Supply Voltage (V)

AD590LF Dev #3

0

5

10

15

20

25

30

Supply Voltage (V)

AD590LF Dev #29

0

50

100

150

200

250

300

0

5

10

15

20

25

30

Supply Voltage (V)

AD590LF Dev #30

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

50

100

150

200

250

300

0

5

10

15

20

25

30

Supply Voltage (V)

AD590LF Dev #31

292K

268K

243K

217K

191K

165K

144K

123K

102K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

292K

268K

243K

217K

191K

165K

144K

123K

102K

80K

292K

268K

243K

217K

191K

165K

144K

123K

102K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K


[image: image2.wmf]0

5

10

15

20

25

30

Supply Voltage (V)

AD590LF Dev #56

0

50

100

150

200

250

300

350

400

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

5

10

15

20

25

30

Supply Voltage (V)

AD590LF Dev #57

0

50

100

150

200

250

300

350

400

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

5

10

15

20

25

30

Supply Voltage (V)

AD590LF Dev #58

0

5

10

15

20

25

30

Supply Voltage (V)

AD590LF Dev #59

0

50

100

150

200

250

300

0

5

10

15

20

25

30

Supply Voltage (V)

AD590LF Dev #60

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

50

100

150

200

250

300

0

5

10

15

20

25

30

Supply Voltage (V)

AD590LF Dev #61

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K


[image: image3.wmf]0

5

10

15

20

25

30

Supply Voltage (V)

AD590JH Dev #10

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JH Dev #11

0

50

100

150

200

250

300

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JH Dev #12

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JH Dev #13

0

50

100

150

200

250

300

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JH Dev #14

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

50

100

150

200

250

300

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JH Dev #15

295K

270K

245K

220K

195K

170K

160K

150K

140K

130K

120K

110K

100K

90K

80K

295K

270K

245K

220K

195K

170K

160K

150K

140K

130K

120K

110K

100K

90K

80K

295K

270K

245K

220K

195K

170K

160K

150K

140K

130K

120K

110K

100K

90K

80K

295K

270K

245K

220K

195K

170K

160K

150K

140K

130K

120K

110K

100K

90K

80K

295K

270K

245K

220K

195K

170K

160K

150K

140K

130K

120K

110K

100K

90K

80K

295K

270K

245K

220K

195K

170K

160K

150K

140K

130K

120K

110K

100K

90K

80K


[image: image4.wmf]0

5

10

15

20

25

30

Supply Voltage (V)

AD590JH Dev #16

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JH Dev #17

0

50

100

150

200

250

300

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JH Dev #18

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JH Dev #19

0

50

100

150

200

250

300

295K

270K

245K

220K

195K

170K

160K

150K

140K

130K

120K

110K

100K

90K

80K

295K

270K

245K

220K

195K

170K

160K

150K

140K

130K

120K

110K

100K

90K

80K

295K

270K

245K

220K

195K

170K

160K

150K

140K

130K

120K

110K

100K

90K

80K

295K

270K

245K

220K

195K

170K

160K

150K

140K

130K

120K

110K

100K

90K

80K


[image: image5.wmf]0

5

10

15

20

25

30

Supply Voltage (V)

AD590JH Harris Dev #20

0

50

100

150

200

250

300

350

400

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JH Harris Dev #21

0

50

100

150

200

250

300

350

400

0

50

100

150

200

250

300

350

400

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JH Harris Dev #22

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JH Harris Dev #23

0

50

100

150

200

250

300

350

400

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JH Harris Dev #24

0

50

100

150

200

250

300

350

400

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

50

100

150

200

250

300

350

400

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JH Harris Dev #25

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K


[image: image6.wmf]0

5

10

15

20

25

30

Supply Voltage (V)

AD590JH Harris Dev #26

0

50

100

150

200

250

300

350

400

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JH Harris Dev #27

0

50

100

150

200

250

300

350

400

0

50

100

150

200

250

300

350

400

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JH Harris Dev #28

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K


[image: image7.wmf]0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #4

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #5

0

50

100

150

200

250

300

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #6

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #7

0

50

100

150

200

250

300

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #8

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

50

100

150

200

250

300

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #9

292K

268K

243K

217K

191K

165K

144K

123K

102K

80K

292K

268K

243K

217K

191K

165K

144K

123K

102K

80K

292K

268K

243K

217K

191K

165K

144K

123K

102K

80K

292K

268K

243K

217K

191K

165K

144K

123K

102K

80K

292K

268K

243K

217K

191K

165K

144K

123K

102K

80K

292K

268K

243K

217K

191K

165K

144K

123K

102K

80K


[image: image8.wmf]0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #32 

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #33

0

50

100

150

200

250

300

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #34

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #35

0

50

100

150

200

250

300

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #36

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

50

100

150

200

250

300

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #37

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K


[image: image9.wmf]0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #38

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #39

0

50

100

150

200

250

300

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #40

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #41

0

50

100

150

200

250

300

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #42

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

50

100

150

200

250

300

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #43

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K


[image: image10.wmf]0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #44

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #45

0

50

100

150

200

250

300

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #46

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #47

0

50

100

150

200

250

300

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #48

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

50

100

150

200

250

300

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #49

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K


[image: image11.wmf]0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #50

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #51

0

50

100

150

200

250

300

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #52

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #53

0

50

100

150

200

250

300

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #54

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

50

100

150

200

250

300

0

5

10

15

20

25

30

Supply Voltage (V)

AD590JF Dev #55

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K

295K

220K

170K

150K

140K

130K

120K

110K

100K

90K

80K


To illustrate the effect of low temperature exposure on the linearity of these temperature transducers, data pertaining to six selected devices are plotted in the following figures.  The data shows the linearity curves for devices 1 and 56 of the LF-series, 10 and 22 of the JH-series, and 6 and 52 of the JF-series.  For a given device, the figure depicts the output current corresponding to a certain temperature at various supply voltage levels.  It can be clearly seen that all devices exhibit some non-linearity as the temperature is decreased. These effects become more apparent at extreme low temperatures and at high supply voltage levels. The severity of these changes are, however, dependent on the series-type of the device, its manufacturer as well as production lot.  

[image: image12.wmf]0

50

100

150

200

250

300

Temperature (K)

AD590LF Dev #1

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

50

100

150

200

250

300

Temperature (K)

AD590LF Dev #56

0

50

100

150

200

250

300

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

50

100

150

200

250

300

Temperature (K)

AD590JH Dev #10

0

50

100

150

200

250

300

Temperature (K)

AD590JH Harris Dev #22

0

50

100

150

200

250

300

0

50

100

150

200

250

300

Temperature (K)

AD590JF Dev #6

0

50

100

150

200

250

300

M

e

a

s

u

r

e

m

e

n

t

 

C

u

r

r

e

n

t

 

(

u

A

)

0

50

100

150

200

250

300

0

50

100

150

200

250

300

Temperature (K)

AD590JF Dev #52

Supply Voltage

30V

10V

5V

4V

Supply Voltage

30V

10V

5V

4V

Supply Voltage

30V

10V

5V

4V

Supply Voltage

30V

10V

5V

4V

Supply Voltage

30V

10V

5V

4V

Supply Voltage

30V

10V

5V

4V


Conclusion


Temperature transducers are widely used in electronic systems for remote sensing applications.  Design simplicity of the AD590 device, which does not require any supporting circuitry, makes it an attractive candidate for use in NASA missions.  A large number of these devices were investigated in this work for suitability for use in low temperature environment. Ceramic-packaged and metal-canned devices, with different accuracy ratings, were included in this investigation to determine if the low temperature exposure has any effect on the packaging material of these components.


Although these devices are rated for only -55 (C (218K), this effort was undertaken to determine their operation limit at low temperature.  All devices exhibited variation in their operational behavior with temperature.  While some of these devices displayed good performance with temperature all the way down to 80K (around liquid nitrogen temperature), others started to deviate in their output response at temperatures below 220K (manufacturer rating).  It is important to note that although the majority of the devices exhibited changes, with varying degrees, in their output linearity with decreasing temperature, none underwent any catastrophic failure due to exposure to the low temperature environment.  This is evident from the full recovery in performance exhibited by all the devices tested upon their thermal stabilization back to room temperature. As far as the packaging material is concerned, the results of this preliminary work preclude any decisive assessment of the effects induced by the low temperature exposure.  More comprehensive testing is, therefore, required to fully characterize the behavior of these devices and to determine their suitability for use and limitation in low temperature environments.  Issues such as thermal cycling, non-linearity, and low temperature operation capability need to be addressed for designing efficient and reliable electronic systems for space and terrestrial applications.

Acknowledgment


This work has been performed at the NASA Glenn Research Center under SETAR and GESS Tasks. Support was also provided by the Jet Propulsion Laboratory through the NASA Electronic Parts and Packaging (NEPP) Program.

References

[1].
M. P. Timko, “A Two-Terminal IC Temperature Transducer”, IEEE Journal of solid-state circuits, Vol. SC-11, No. 6, December 1976.

[2].
Analog Devices, Inc. product data sheet.

[3].
Harris/Intersil Corporation product data sheet.

16
NASA GRC 8/01


