COTS PEMS Packaging Document  2 of 2

           




NEPP Program

[image: image1.wmf]








[image: image2.png]












   
Package Qualification Testing Results
Performed on the 5 Part Types in the NEPP/NEPAG PEMS Study 

FY ’03 Work

Author:

R. David Gerke







Jet Propulsion Laboratory







david.gerke@jpl.nasa.gov






818-393-6372

Reviewed By: 
Jeannette Plante
Dynamic Range Corp. (GSFC)



jfplante@pop500.gsfc.nasa.gov



301-286-7437

http://nepp.nasa.gov
Date: February 2004

1. Executive Summary TC "Executive Summary" \f C \l "1" 
 A group each of the five NEPP/NEPAG parts were subjected to Qualification Testing (Temperature Cycling and HAST). Prior to the Qualification Testing each part was preconditioned (to simulate the assembly process) to JESD22-A113C “Preconditioning of Nonhermetic Surface Mount Devices Prior to Reliability Testing”.  The temperature cycle testing was conducted according to MIL-STD 883 condition C (-65 oC to + 150 oC) for 1000 cycles. There were two (2) read points; one at 500 cycles and the other at 1000 cycles. The HAST test was conducted in accordance of JESD22-A110-B which states the following conditions: 130 ± 2 oC, 85 % ± 5 RH, and 33.3 psia for 96 +/- 2 hours.
All parts were electrically tested at Room Temperature (RT) to the manufacturer’s data sheet with testing for both gross and parametric functionality. The test houses kept control units for comparison to tested devices. Also, all parts were tested in the “as received” condition from the vendor with no additional screening or testing (with the exception of initial electrical testing to insure that each part is electrically good prior to the qualification flow testing).
The purpose of this evaluation is to take the parts in the “as-received” condition from each vendor with no additional screening or testing and subject them to preconditioning and qualification stresses that are similar to those used by the manufacturer in their qualification testing, the results of which are reported to customers. The results from this evaluation, on as-received unscreened parts, will be compared and contrasted to future (FY ’04) work that will be performed on the same part types that have been through a screening flow. The premise behind this study is that all parts should go through this preconditioning and qualification stress flow and pass with no failures with parts straight from the vendor as the vendors of the 5 part types claim that their parts can withstand these types of stresses.  The Moisture Sensitivity Level Evaluation (previously reported – Part 1 of 2) showed that not all of the part types met the MSL rating that was quoted by their respective manufacturer.
There were failures resulting from the qualification stresses in 2 part types, one set of failures in temperature cycle testing and the other set of failures were in the HAST testing. Also, there was a functional failure during the preconditioning stage of the flow. The results will be summarized by assigning a level of concern as indicated below for each step of the testing: 
· The Qualification Test Results including;

· Precondition overall assessment,               low       high
· Temperature Cycle Results, and                low       high
· HAST Results,                                           low       high
· The final page contains the Overall scoring for each Part Type with some reference to the MSL testing (report 1 of 2 reported in February 2004) and some conclusions/observations.

2. Evaluation Details TC "Evaluation Details" \f C \l "1" 
a. Introduction TC "Introduction" \f C \l "2" 
The JESD22-A113C Test Method establishes an industry standard preconditioning flow for non-hermetic solid state SMDs (surface mount devices) that is representative of a typical industry multiple solder reflow operation. SMDs are subjected to the appropriate preconditioning sequence of the specification by the semiconductor manufacturer prior to being submitted to specific in-house reliability testing (qualification and reliability monitoring) to evaluate long term reliability (which might be impacted by solder reflow). The reliability numbers published by the manufacturers reflect a composite of stresses which include those from both the preconditioning test flow and the qualification test flow. This Qualification Validation is one portion of the overall packaging-related test plan which is a subset evaluation of the larger NEPP/NEPAG PEMS study. 

b. Preconditioning:  Test Conditions and Results
 TC "Test Conditions" \f C \l "2" 
In order to subject the parts to the proper precondition, the vendors of the 5 different part types had to disclose the moisture sensitivity level (MSL) to which they have qualified their respective PEMS devices.  A memo submitted to the NEPP/NEPAG team, dated November 20, 2002 and authored by Jeannette Plante, summarized each respective vendor’s quoted moisture sensitivity levels of each of the five part types. The information was obtained by website review as well as email correspondence with Quality Assurance Personnel from the respective manufacturers, as necessary. Two of the part types are manufactured with lead-free plating on the leads and were stated as able to pass the test condition simulating a higher solder reflow temperature. The JESD22-A113 Specification covers the differences between peak reflow-temperatures for tin-lead and lead-free solders.  For small body devices (which these parts qualify as) the peak solder reflow temperatures are: tin-lead solder ( 240 +0/-5°C  and the lead-free solder ( 250 +0/-5°C (temperatures are to be measured at the topside of the part).  The reflow method chosen for this evaluation was convection reflow as it is the preferred method called out by JESD22-A113 and it is the most representative reflow method used by volume assembly manufacturers. Since this is a “Package” Evaluation, room temperature electrical tests were performed (per JESD22-A113C).  A more in-depth look at the Packaging Stresses effect on the Electrical Performance of these devices will follow in FY’04.

The evaluated parts are listed below in Table 1 with their reported MSL rating:

Table 1: Test Samples and Their Manufacturer MSL and Lead Finish Ratings.

	Part Type
	Sample
Size*
	Vendor
	MSL
	Moisture Condition

Time (hours)             Conditions
	Lead Finish
	Reflow Temperature**

	A/D
	24
	A
	1
	168 (+5/-0)
	85 °C/85% RH
	Lead-free
	250 +0/-5°C

	Multiplexer
	24
	B
	1
	168 (+5/-0)
	85 °C/85% RH
	Tin-lead
	240 +0/-5°C

	Op Amp
	24
	C
	1
	168 (+5/-0)
	85 °C/85% RH
	Tin-lead
	240 +0/-5°C

	Reference
	24
	D
	1
	168 (+5/-0)
	85 °C/85% RH
	Lead-free
	250 +0/-5°C

	Amplifier
	24
	E
	3
	192 (+5/-0)
	30 °C/60% RH
	Tin-lead
	240 +0/-5°C


*   Sample size of 22 units were needed, to be split into 2 samples of 11 each; there were 2 spares
** Per JEDEC JESD22-A113C Specification flow
A sample of 24 units for each of the 5 part types were selected and tested to the JEDEC JESD22-A113C Specification test flow as follows:

1.  Serialization, added for the NEPP/NEPAG investigation 

2.  Initial Electrical Test, Test appropriate electrical parameters, e.g., data sheet values. Replace any devices that fail to meet          

     this requirement.
3.  Initial Inspection, 40X visual (required)

4.
 CSAM (C-Mode Scanning Acoustic Microscopy), added for the NEPP/NEPAG investigation.
5.  Temperature Cycling, (-40 +0/-10) °C to (60 °C +10/-0) °C per JESD22-A104 to simulate the shipping of the devices from   

     the manufacturers to distributors, etc.
6.  Bake, 24 hours minimum at 125 +5/-0°C.
7. 
Moisture Soak, Level 1 ( 85oC/85%RH for 168 hours +5/-0, Level 3 ( 30oC/60%RH for 192 hours +5/-0.
8. 
Reflow, Not sooner than 15 minutes and not longer than 4 hours after removal from the temperature/humidity chamber, subject the sample to 3 cycles of the appropriate reflow conditions. For details see Appendix A.
9.   Flux Application, After the reflow solder cycles are completed, allow the devices to cool at room ambient for 15

      minutes minimum. Apply an activated water soluble flux to the device leads by bulk immersion of the entire parts in flux at    

      room ambient for 10 seconds minimum.
10. Cleaning, Clean devices externally using multiple agitated deionized water rinses. No waiting time is required between flux   

      application and cleaning.
11. Drying, Devices should be dried at room ambient prior to submission to reliability testing.
11.
Final Electrical Test, Test appropriate electrical parameters, e.g., data sheet values.
12.
Final Acoustic Microscopy, added for the NEPP/NEPAG investigation.

The details of JEDEC JESD22-A113C and the reflow profiles used in this evaluation are presented in Appendix A.  The results of this part of the evaluation are presented next in Table 2. There was one failure out of a total of 120 devices (5 part types x 24 devices/part type). The device failure was a functional reject, a hard failure, for which the cause of the failure was not determined. In JESD22-A113C, an electrical failure at the end of the flow could be attributed to the part type failing the MSL level unless otherwise proven by failure analysis.  From the NEPP/NEPAG report entitled “Moisture Sensitivity Level (MSL) Packaging Evaluation” it was documented that the A/D part from Vendor A indeed did not pass the MSL level 1 that was claimed by Vendor A.  It was beyond the scope of this evaluation to investigate the devices at a lower MSL – one that was perhaps more appropriate for the device. The solution for this issue was to use one of the spare devices and substitute it for the failed device and continue the testing.  
Table 2: Electrical Performance Summary following the Pre-conditioning Flow
	Part Type
	Sample

Size*
	Vendor
	MSL
	Lead Finish
	Die

Coat
	Electrical Test (RT)
(Rejects/Total)

   Functional             Parametric
	Final Assessment

	A/D
	24
	A
	1
	Pb-free
	no
	1/24
	0/24
	fail

	Multiplexer
	24
	B
	1
	Tin-lead
	no
	0/24
	0/24
	pass

	Op Amp
	24
	C
	1
	Tin-lead
	yes
	0/24
	0/24
	pass

	Reference
	24
	D
	1
	Pb-free
	yes
	0/24
	0/24
	pass

	Amplifier
	24
	E
	3
	Tin-lead
	no
	0/24
	0/24
	pass


* Total number of parts required was 22, 2 parts were spares.

As can be seen in Table 2, one part exhibited a functional failure. The A/D part from Vendor A, failed the functional test following the preconditioning step. A spare part was substituted for the functional failed A/D part that the evaluation could continue with a full set of devices. The cause of the failure is not known at this time.

c. Temperature Cycling:  Test Conditions and Results
The temperature cycle testing was conducted according to MIL-STD 883 condition C (-65 oC to + 150 oC) for 1000 cycles. There were two (2) read points; one at 500 cycles and the other at 1000 cycles. The Reference part from Vendor D had failures in both read points of temperature cycles. There were 7 failures after 500 cycles and all devices failed at the 1000 cycle read point. The results of the temperature cycle testing are shown below in Table 3. As of the writing of this report, no F/A has been conducted on these parts. Failure analyses of parametric failures are very difficult to resolve because the part is still functioning but has drifted out of the manufacturer’s specification limits. This part type’s parameter that drifted beyond the manufacturer’s limits was VOS.
Table 3: Overview of Temperature Cycle Results

	Part Type
	Vendor
	500 T/C Read Point

Electrical Test (RT)

(Rejects/Total)

Functional    Parametric
	1000 T/C Read Point

Electrical Test (RT)
(Rejects/Total)

Functional    Parametric
	Final Assessment

	A/D
	A
	0/11
	0/11
	0/11
	0/11
	pass

	Multiplexer
	B
	0/11
	0/11
	0/11
	0/11
	pass

	Op Amp
	C
	0/11
	0/11
	0/11
	0/11
	pass

	Reference
	D
	0/11
	7/11
	0/11
	11/11
	fail

	Amplifier
	E
	0/11
	0/11
	0/11
	0/11
	pass


d. HAST:  Test Conditions and Results
The HAST test was conducted in accordance of JESD22-A110-B which states the following conditions: 130 ± 2 oC, 85 % ± 5 RH, and 33.3 psia for 96 +/- 2 hours. The Op Amp from Vendor C exhibited 11 failures out of 11 possible devices following the HAST test.  The parameter that drifted beyond the manufacturer’s specification limits was Vout. A quick destructive physical analysis (DPA) was conducted on one of the failures looking for signs of corrosion. No signs of corrosion were found. So, it is thought that corrosion is not the mechanism that is shifting the parameter.  Like the parametric failures in the temperature cycle testing, the cause is a very difficult thing to determine.
Table 4: HAST Results

	Part Type
	Vendor
	96 Hrs. HAST Read Point

Electrical Test (RT)
(Rejects/Total)

Functional      Parametric
	Final Assessment

	A/D
	A
	0/11
	0/11
	pass

	Multiplexer
	B
	0/11
	0/11
	pass

	Op Amp
	C
	0/11
	11/11
	fail

	Reference
	D
	0/11
	0/11
	pass

	Amplifier
	E
	0/11
	0/11
	pass


3. Discussion and Conclusions
The results in Tables 3 and 4 suggest that some of the parts performed as was stated by the manufacturers and some did not.  Since each part type represents a different vendor and wafer fab process, there are no similarities that stand out as obvious contributors to the shifts in performance.  However it is interesting to note that both non-compliant part types did have a die coating placed on at assembly prior to the molding process. Generally, the addition of a die coat is thought of as a performance enhancer, and in these cases was likely used as an assembly yield enhancer but in this case, it did not perform in that way.  In addition to the parametric failures, there was also a functional failure from Vendor A following the preconditioning stressing. All of these failures suggest that the use of PEMS requires quality and reliability validation through screening and qualification testing performed by the user, prior to using the parts.

In order for JESD22-A113C to be completed to the “letter of the law”, the IPC/JEDEC J-STD-020B Specification entitled “Moisture/Reflow Sensitivity Classification for Nonhermetic Solid State Surface Mount Devices” should be evaluated for the devices under test (DUT). The NEPP/NEPAG Report entitled “Moisture Sensitivity Level (MSL) Packaging Evaluation” submitted in February 2004 covers that area of investigation.  In that report, the MSL testing showed that Vendors A, B and C did not pass the Level 1 rating claimed by the manufacturers as those 3 part types did not pass the CSAM inspection portion of the test. Even though those parts did not pass the CSAM inspection portion of the test, all of the parts did pass the electrical test portion. It was decided to use the stated MSL levels for this evaluation to investigate the possible ramifications of trusting the vendors and handling the parts per their instructions and study parts in the “as-received” condition. As stated earlier, Packaging Qualification Work in FY’04 will study the effects of screening on the results of the pre-conditioning flow.
The results of the two Packaging Evaluations performed (in FY’03) on the 5 part types selected by the NEPP/NEPAG team suggest that a carefully thought-out screening flow should be employed on PEMS as only 2 part types out of 5 total part types passed all of the Qualification Level testing (this report).  Also in the previously reported work entitled “Moisture Sensitivity Level (MSL) Packaging Evaluation” submitted in February 2004, only 2 part types passed the MSL testing. Table 5 shows the roll-up summary of the MSL testing and the Package Qualification Testing performed so far. 
Table 5. Roll-up Table of all Packaging Test Results

Electrical Test results include Functional (F) and Parametric (P).
	
	
	
	MSL Evaluation
(Report 1 of 2)
	Qualification Testing
(Report 2 of 2 – this report)
	     Final

Assessment

	Part Type
	Vendor
	MSL

Vendor

Rating
	Overall


	Preconditioning*


	HAST
	Temperature      

     Cycle
	        All

     Testing

	A/D
	A
	1
	fail
	fail (F)
	pass
	pass
	fail

	Multiplexer
	B
	1
	fail
	pass
	pass
	pass
	fail

	Op Amp
	C
	1
	fail
	pass
	fail (P)
	pass
	fail

	Reference
	D
	1
	pass
	pass
	pass
	fail (P)
	fail

	Amplifier
	E
	3
	pass
	pass
	pass
	pass
	pass


* Note: CSAM inspection was also performed as an additional step and the results were identical to the MSL CSAM inspections (i.e. the same part types failed). It is not a requirement of preconditioning to perform CSAM inspection; therefore it is not reported here.

 TC "Results" \f C \l "1" 
As can be seen in Table 5, only 1 part type passed all of the package level testing; the Amplifier manufactured by Vendor E. That is, the Amplifier Part has a pass (or a green box) all the way across in Table 5. Every other part type failed at least one test or precondition. It is noteworthy that Vendor E elected to rate their part type as an MSL of Level 3 rather than push for a higher rating of 2.  All of the other vendors rated their parts as MSL Level 1, the highest level available - nearly equivalent to hermetic parts with regards to environmental exposure to moisture on the production floor. A large number of CSAM delamination failures were noted throughout both packaging evaluations for the MSL Level 1 parts.  Since it was beyond the scope of these evaluations to investigate the reliability of the parts should an MSL Level 2 be specified by the vendors, no further work was performed.  However,  given the fact that Vendor E with their recommendation of MSL Level 3, was the only part type to pass all testing, it could be concluded that the vendors of the other part types were over zealous in their rating their respective part types as MSL Level 1.  It is noteworthy that Vendor B failed CSAM inspection but passed all other electrical performance testing, including all qualification testing. For this part, there was no strong correlation between CSAM inspection and reliability (T/C and HAST).

Nearly all electrical failures (with the exception of the precondition failure of Vendor A) were parametric failures.  The parts functioned but still failed the data sheet specifications. It is hoped that the FY’04 package level testing on screened parts would screen out weak parts and pass similar Packaging Level testing. 
Appendix A. TC "Appendix A." \f C \l "1" 
a. Reflow Conditions: TC "Reflow Conditions:" \f C \l "2"  

The conditions from the JESD22-A113C Specification state that the solder reflow occur not sooner than 15 minutes and not longer than 4 hours after removal from the temperature/humidity chamber, subject the sample to 3 cycles of the appropriate reflow conditions as defined in Table A-1 and Figure A-1. If the timing between removal from the temperature/humidity chamber and initial reflow cannot be met then the parts must be rebaked and resoaked . The time between reflows shall be 5 minutes minimum and 60 minutes maximum.

Table A-1: Reflow conditions from the JESD22-A113C Specification

[image: image3.png]Reflow condition

Sn-Pb eutectic assembly

Pb-Free assembly

Pkg
thickness
225mmor
Pke. volume
2350 mm’

Phg
thickness
<2.5mmand
Pkg. volume

0 mm*

Pkg
thickness
225mmor
Pke. volume
2350 mm’

Phg
thickness
<2.5mmand
Pkg. volume
50 mm*

Average ramp-up rate
(Liquidus Temperature (Ty) to
Peak)

3 “Clsecond max.

3 “Clsecond max.

Preheat
- Temperature Min (T.(uy)
- Temperature Max (Tumas)
- Time (in to max) (&)

100°C
150°C
60-120 seconds

150°C
200°C
60-180 seconds

e 10 Te
- Ramp-up Rate

3 *Clsecond max.

Time maintained above:
- Temperature (Ty)
- Time(t)

183°C
60-150 seconds

217°C
60-150 seconds

Peak Temperature (T,)

240 +0/-5 °C

245+0/-5°C

0+0/-5 °C

Time within 5 °C of actual Peak
Temperature (t;)

10-30 seconds

10-30 seconds

10-30 seconds

2040 seconds

Ramp-down Rate

6 °Clsecond max.

6 °Clsecond max.

Time 25 °C to Peak Temperature

6 minutes max.

8 minutes max.

NOTE  All temperatures refer to topside of the package, measured on the package body surface.





[image: image4.png]Temperature =)

B

el

Ramp-up

4l Zong]

Time =)





Figure A-1: Graphic representation of the reflow conditions from the IPC JESD22-A113C Specification

b. Profiles used in this Evaluation TC "Profiles used in this Evaluation" \f C \l "2" 
The convection reflow profiles were performed in a 5 zone Heller (100%) convection (production worthy) furnace. Particular care was exercised for furnace loading and to simulate the conditions used for the evaluation parts. The parts were ‘reflowed’ on PWB material or pallets and a thermocouple was mounted on top of similar parts to obtain the correct profiles.  JESD22-A113C specifies that the temperature is to be measured at the top of the parts as the heat contacts the outside surfaces prior to penetrating to the inside of the part.

Representative profiles used in this evaluation are shown below in Figures A-3 and A-4 for the tin-lead and lead-free parts, respectively.

	[image: image5.png]le 1500.KOF ™
ata Time Thu 05/22/03 15:30:15 .

art Time Thu 06/22103 15:21:54 Slim K | C

/g Qven Temp - .

acipe TOPLINE.KPF Transmitter/Data Logger

roduct CAWINKIC\PRODUCTS\TOPLINE. KBF

| 55 i Belt Speed 36.0 cm/min

317

242

Rdednlh b

92

(8]
(o))

0.0 186 2.4 3.2 4.0 4.8 6.4

Minutes (slope Degrees/Second)







Figure A-2: Furnace profile used for parts with tin-lead lead-finish. The peak top-side component temperature was 238.1 oC which meets the specification for small bodied components of between 235 oC and 240 oC.
	[image: image6.png]File 1500.KOF ™
Profiling Live Thu 05/29/03 11:57:25 H

Start Time Thu 05/29/03 11:47:39 Slim KI C
Avg Oven Temp Transmitter/Data Logger

Recipe PBFRESB.KPF
Product CAWINKIC\PRODUCTS\TOPLINE. KBF

55 { Belt Speed 30.0 cm/min

323

248 =S

Celsius
\

|4

29 2 3 4 5 6 7 8
| Minutes (slope Degrees/Second)






Figure A-3: Furnace profile used for parts with tin-lead lead-finish. The peak top-side component temperature was 248.9 oC which meets the specification for small bodied components of between 245 oC and 250 oC.
R





GG





G





GG





R





GG





G





GG





GG





R





GG





R





G





GG





G





GG





G





GG





G





GG





G





GG





R





GG





R





GG





G





GG





R





GG





G





GG





G





GG





G











R





GG





GG





G











G





GG





GG





G











GG











GG





G



























































































































































































































































































































































G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





G





GG





R





GG





R





GG





G





GG





G





GG





G











G





GG





G





GG





G





GG





G





GG





GG





R





GG





R





GG





G





GG





G





GG





R





GG





R





GG





R





GG





G





GG





G





GG





G





GG





R





GG





R





GG





G





GG





R





GG





G





GG





G





GG





G





GG





G





GG





R





GG





R





GG





G





GG





G





GG





G





GG





G





GG





G








8

